Supplementary data 1 – Sheep ELISA ROC Analysis

Serum samples were collected from 250 sheep that had matched individual FEC or from slaughter at a local abattoir (North West England). Sheep were confirmed positive or negative for liver fluke by liver inspection.

All samples were run on the UoL Sheep ELISA, following the protocol described in Walsh et al 2021, and results expressed as a percentage of the positive control (PP).

ELISA PP results for each sample were paired with true status based on FEC or abattoir data. Data was analysed using R (4.3.3) using RStudio (version 2024.12.1+563) using the pROC function (Robin et al 2011). The function was allowed to suggested the PP threshold for optimum Se and Sp.

Results

Results of the pROC function are shown in Table 1. The function indicated that a PP of 24 was an optimum threshold of the cut off. At this threshold, the Se was 87% and the Sp 76%.

Table 1 – Outputs from pROC analysis. The optimum cut off is hightlighted

| thresholds | sp.low | sp.median | sp.high | se.low | se.median | se.high |
| --- | --- | --- | --- | --- | --- | --- |
| 6.9832 | 0.3423 | 0.4324 | 0.5225 | 1 | 1 | 1 |
| 12.87875 | 0.4865 | 0.5856 | 0.6757 | 0.9784 | 0.9928 | 1 |
| 13.82525 | 0.5135 | 0.6126 | 0.7027 | 0.964 | 0.9856 | 1 |
| 14.61475 | 0.5403 | 0.6306 | 0.7207 | 0.9353 | 0.964 | 0.9928 |
| 15.20525 | 0.5495 | 0.6396 | 0.7207 | 0.9209 | 0.9568 | 0.9856 |
| 17.21694 | 0.5586 | 0.6486 | 0.7387 | 0.8921 | 0.9353 | 0.9712 |
| 18.094 | 0.6036 | 0.6937 | 0.7748 | 0.8849 | 0.9281 | 0.9712 |
| 19.38051 | 0.6124 | 0.7027 | 0.7838 | 0.8633 | 0.9137 | 0.9568 |
| 19.99005 | 0.6216 | 0.7117 | 0.7928 | 0.8561 | 0.9065 | 0.9496 |
| 21.23625 | 0.6394 | 0.7207 | 0.793 | 0.8417 | 0.8921 | 0.9424 |
| 22.512 | 0.6396 | 0.7297 | 0.8018 | 0.8201 | 0.8777 | 0.9281 |
| 23.6235 | 0.6757 | 0.7568 | 0.8288 | 0.8129 | 0.8705 | 0.9209 |
| 28.24318 | 0.6937 | 0.7748 | 0.8468 | 0.741 | 0.8058 | 0.8705 |
| 30.33241 | 0.7207 | 0.8018 | 0.8739 | 0.6978 | 0.7698 | 0.8345 |
| 34.8515 | 0.7568 | 0.8288 | 0.8919 | 0.6906 | 0.7626 | 0.8273 |
| 36.884 | 0.7658 | 0.8378 | 0.9009 | 0.6619 | 0.7338 | 0.8058 |
| 39.2225 | 0.7748 | 0.8468 | 0.9099 | 0.6043 | 0.6835 | 0.7554 |
| 42.13921 | 0.7838 | 0.8559 | 0.9189 | 0.5612 | 0.6403 | 0.7194 |
| 43.4455 | 0.7928 | 0.8649 | 0.9279 | 0.5252 | 0.6043 | 0.6906 |
| 44.538 | 0.8108 | 0.8739 | 0.9369 | 0.5108 | 0.5899 | 0.6763 |
| 51.002 | 0.8198 | 0.8829 | 0.9369 | 0.4676 | 0.5468 | 0.6331 |
| 54.40975 | 0.8468 | 0.9099 | 0.964 | 0.446 | 0.5252 | 0.6115 |
| 56.39959 | 0.8559 | 0.9189 | 0.964 | 0.4245 | 0.5036 | 0.5899 |
| 61.56425 | 0.8739 | 0.9279 | 0.973 | 0.3813 | 0.4604 | 0.5468 |
| 61.97787 | 0.8919 | 0.9369 | 0.982 | 0.3741 | 0.4532 | 0.5396 |
| 63.92925 | 0.9009 | 0.9459 | 0.982 | 0.3597 | 0.4388 | 0.5252 |
| 69.10975 | 0.9099 | 0.955 | 0.991 | 0.3237 | 0.4029 | 0.482 |
| 71.3445 | 0.9279 | 0.964 | 0.991 | 0.2948 | 0.3669 | 0.446 |
| 95.68144 | 0.955 | 0.982 | 1 | 0.1367 | 0.2014 | 0.2734 |
| 104.3541 | 0.973 | 0.991 | 1 | 0.07914 | 0.1367 | 0.1942 |
| 119.0346 | 1 | 1 | 1 | 0.05036 | 0.09353 | 0.1439 |

Results were then plotted by PP value colour coded by their positive (blue) and negative (pink) status (Figure 1). The purple line represents a cut off of 24 PP and shows that at this cut off, the ELISA result is matching the true positive status.

Figure 1 – bar chart to show number of individuals testing at each PP value and coded via their true positive or negative status.

Whilst most animals above 24 PP are true positives, there is a proportion that showed as true negatives. Antibodies are known to persist for several weeks following infection (Salimi-Bejestani 2007). Therefore, if animals had been treated in the weeks prior to sampling, this could create a false positive result by ELISA. It is also difficult to observe early infection using liver inspection, therefore could lead to false negatives by this method (Rapsch et al 2006, Mazeri et al 2016). Finally, the Se and Sp of FEC can also be variable, due to it only confirming infection when adults are present at 8-12 weeks post infection and eggs can become trapped in the gall bladder following treatment, again leading to false results (Fairweather et al 2011, Charlier et al 2014, Mazeri et al 2016, Rapsch et al 2006).

References

Fairweather (2011) Reducing the future threat from (liver) fluke: realistic prospect or quixotic fantasy? Veterinary Parasitology. 180 (1-2), 133-143

Mazeri et al (2016) Evaluation of the performance of five diagnostic tests for Fasciola hepatica infection in naturally infected cattle using a Bayesian no gold standard approach. PLoS One 11 (8), e0161621

Robin et al (2011) pROC: an open source package for R and S+ to analyze and compare ROC curves. BMC Bioinformatics 12, 77, doi.org/10.1186/1471-2105-12-77

Rapsch et al (2006) Estimating the true prevalence of Fasciola hepatica in cattle slaughtered in Switzerland in the absence of an absolute diagnostic test. International Journal for Parasitology 36 (10-11) 1153-1158

Salimi-Bejestani et al (2007) Evaluation of an enzyme-linked immunosorbent assay for detection of antibodies to Fasciola hepatica in milk. Veterinary Parasitology. 149 (3-4) 290-293